## Sheet1
Key Allegations | Event Description | Confirmation | Updated Event Description
1. Out of box anomaly/packaging/labeling issue 2. Temperature indicator activated | It was reported that an AVALUS™ ULTRA BIOPROSTHESIS 25MM VALVE was received with an out of box anomaly related to packaging and labeling, specifically an activated temperature indicator. The device was never implanted and was in the possession of the hospital. The issue will be resolved by sending a replacement. There was no patient involved. | {'Correct'} | No Update Needed
1. P-wave oversensing in the Implantable Cardioverter Defibrillator. 2. More than 6 seconds of Tidal Volume-NS episodes within an hour. 3. VF (Ventricular Fibrillation) count greater than halfway to VF Number of Beats to detect. 4. Increased susceptibility to Electromagnetic Interference with unipolar setting. 5. Allegations or indication of sensing issues. | It was reported that there were sensing issues with the Aurora EV-ICD, specifically P-wave oversensing. The event included more than 6 seconds of Tidal Volume-NS episodes within an hour and a VF count greater than halfway to the VF Number of Beats to detect. There was also an increased susceptibility to Electromagnetic Interference with the unipolar setting. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. Fever 2. Redness | It was reported that there was an issue involving the LEAD 6935M55 DF4 ACTIVE SC SPRINT 55 CM, ICD-DR DDMB1D4 EVERA MRI XT IS-1/DF4 US, and LEAD 5076-45 LEAD BI SIL STEROID SCREW implantable devices. The patient exhibited symptoms of fever and redness. All devices were explanted. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. Possible RV oversensing/undersensing noted on stored electrogram. | It was reported that possible right ventricular (RV) oversensing and undersensing were noted on stored electrograms. No programming changes were requested at the time of the event. It was unknown if there was any patient involvement or complications as a result of the issue. | {'Correct'} | No Update Needed
1. Difficulties in screwing the electrode. 2. Decision to remove, test, and confirm the failure of the electrode. 3. Defective screw connection. | It was reported that during the implant of a SPRINT QUAD RV/SVC electrode, there were difficulties in screwing the electrode. The physician decided to remove the electrode, test it, and confirmed the failure due to a defective screw connection. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. Device withheld for noise. 2. Noise troubleshooting required. 3. Oversensing protection set at 6. 4. Sensitivity set at 0.15 millivolts (mV). 5. Defibrillation threshold tests (DFTs) conducted at 0.45 millivolts (mV). 6. Noise created by device. | It was reported that a defibrillator device created noise, leading to the device being withheld. Noise troubleshooting was required, and the oversensing protection was set at 6. The sensitivity was set at 0.15 millivolts (mV), and defibrillation threshold tests (DFTs) were conducted at 0.45 millivolts (mV). The sense polarity of R1-R2 and the NID, currently at 30/40, were discussed as possible areas to program to address oversensing. The caller was advised to reach out to a local representative for live troubleshooting with the patient in the office. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. Tetralogy of Fallot 2. Surgical Pulmonary Valve Replacement 3. NYHA Class III 4. Fixed leaflet mobility (noted in 1 Year core lab echocardiogram) 5. Leaflet calcification (noted in 1 Year core lab echocardiogram) | It was reported that during a follow-up for a patient with a history of Tetralogy of Fallot and Surgical Pulmonary Valve Replacement, a 1-year core lab echocardiogram noted fixed leaflet mobility and leaflet calcification in the Harmony Transcatheter Pulmonary Valve (TPV). The patient was classified as NYHA Class III. The sponsor considered these findings to be causally related to the study valve. It was unknown if there was any patient involvement or complications as a result of the issue. | {'Correct'} | No Update Needed
1. Severe aortic stenosis 2. Severe calcification 3. Heavy calcification of the valve requiring multiple attempts to cross the annular plane with the wire and delivery system 4. Two different valvuloplasties performed | It was reported that during a transcatheter heart valve procedure, there was severe aortic stenosis and severe calcification, which led to heavy calcification of the valve. This required multiple attempts to cross the annular plane with the wire and delivery system. Two different valvuloplasties were performed. Due to the heavy calcification, several attempts were made to cross the valve with the initial delivery system, resulting in a "mark" on the capsule. Consequently, the delivery system was replaced with a new one, after which the procedure proceeded normally. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. Aortic valve stenosis. 2. Second heart valve implantation due to pop-out of the first valve. 3. The nose cone of the delivery catheter system got caught on the inflow cone of the valve, causing the valve to be pulled into the aorta. 4. Use of a Sapien 29 millimeter valve as a second valve. | It was reported that during a second heart valve implantation procedure, the nose cone of the delivery catheter system got caught on the inflow cone of the valve, causing the valve to be pulled into the aorta. The patient initially had aortic valve stenosis and underwent a procedure involving the implantation of a 34mm EVOLUT FX+ valve via a transfemoral approach, with pre-dilatation using a 24mm Osypka balloon. The valve was loaded into the delivery system according to best practices and positioned under fluoroscopy. After the final release, the valve was stable with an implantation depth of 3mm. However, upon removal of the delivery catheter system, the issue occurred. The physician then decided to use a Sapien 29 millimeter valve as a second valve. The patient remained stable throughout the procedure. | {'Correct'} | No Update Needed
1. Complete heart block after aortic valve replacement. 2. Inability to walk or talk post-procedure. 3. Required rehabilitation and extended hospital stay (close to 3 weeks) following the procedure. 4. Digression in condition requiring additional rehabilitation after a possible transient ischemic attack (TIA). 5. Concern about potential plaque dislodgement causing kidney problems. 6. Two other patients with similar mental and physical outcomes post-procedure. | It was reported that after an Evolut FX+ Transcatheter Aortic Valve Replacement, the patient experienced a complete heart block, necessitating a subsequent pacemaker implant. Post-procedure, the patient was unable to walk or talk and required rehabilitation, resulting in an extended hospital stay of close to three weeks. There was a digression in the patient's condition, requiring additional rehabilitation after a possible transient ischemic attack (TIA). There was also concern about potential plaque dislodgement causing kidney problems. Two other patients reportedly had similar mental and physical outcomes following the same procedure. | {'Correct'} | No Update Needed
1. Severe aortic stenosis (AS) 2. Planned treatment: Transcatheter aortic valve replacement (TAV in TAV) with 26 Sapien valve 3. Device-related issue: Medtronic device (valve) contributed to symptoms according to the physician | It was reported that a patient with severe aortic stenosis (AS) was planned to undergo a transcatheter aortic valve replacement (TAV in TAV) with a 26 Sapien valve. The Medtronic device, specifically the valve, was noted by the physician to have contributed to the symptoms. The patient required an intervention as a result of the event. The device was inspected prior to use, and there was no delay in the procedure. | {'Correct'} | No Update Needed
1. Medtronic transcatheter aortic valve replacement implanted. 2. Medtronic pacemaker implanted within one year. | It was reported that a Medtronic transcatheter aortic valve replacement was implanted, followed by the implantation of a Medtronic pacemaker within one year. No patient complications have been reported as a result of this event. | {'Correct'} | No Update Needed
1. No Medtronic person present at the event. 2. No association with the patient. 3. No replacement of the product by Medtronic. 4. Batteries loaded in 2023, considered too long ago. 5. Customer notified to return the product. | It was reported that the event involved a Power Pack with batteries loaded in 2023, which was considered too long ago. The customer was notified to return the product. There was no patient involved. | {'Correct'} | No Update Needed
1. Device Failure: Short Self Test (SST) failure. 2. Allegation: Customer communicated dissatisfaction with the product. | It was reported that there was a device failure involving a Short Self Test (SST) failure. The customer communicated dissatisfaction with the product. There was no patient involved. | {'Correct'} | No Update Needed
1. Reason for Right Ventricular (RV) lead/Return Repair: Not specified. 2. Fault Reported: BLANK. 3. Allegation of patient death or serious injury: No. 4. Customer communicated dissatisfaction with the product: Yes. | It was reported that during a pre-check, there was dissatisfaction communicated with the product due to a blank screen issue associated with the equipment. No patient complications have been reported as a result of this event. | {'<<<Missing FLAG>>>': ['Reason for Right Ventricular (RV) lead/Return Repair: Not specified.']} | It was reported that during a pre-check, there was dissatisfaction communicated with the product due to a blank screen issue associated with the equipment. The reason for Right Ventricular (RV) lead/Return Repair was not specified. No patient complications have been reported as a result of this event.
1. Critical care unit installation occurred. 2. Exhalation valve (EV) issue. | It was reported that during the installation of a critical care unit, there was an issue with the exhalation valve (EV). There was no patient involved. | {'Correct'} | No Update Needed
1. Power supply malfunction (Damage Code H13). | It was reported that there was a power supply malfunction. There was no patient involved. | {'Correct'} | No Update Needed